# 1 Portable CMM Interface: Faro, Romer, Sandia NL

The SpatialAnalyzer Portable CMM (Coordinate Measuring Machine) Interface application is used to interface to one of three families of Portable CMM’s (also called arms).  The interface communicates to SpatialAnalyzer using New River Kinematics’ TCP/IP socket protocol and communicates to the given CMM arm using the manufacturer’s communication protocol.  Faro.exe and Romer.exe use serial communication.  The Sandia National Labs interface, SandiaArm.exe, uses an interface to a Galil motion control board.  For the Romer and Faro arm interfaces, the buttons on the arms are used in combination with the graphical user interface to run the arms.  Specific keyboard keys are used with the Sandia National Labs arm in place of arm buttons.  The arms are quite different at the instrument level, but the interface applications look and feel identical to the user.

To run any of the arm interfaces, you must first install the arm specific configuration, or parameter file.  This file is available from the given arm manufacturer, and in fact comes with the arm in all cases.  Some manufacturers provide programs which load the files to the arm controllers.

You must also have SpatialAnalyzer installed.  The SA install includes all three arm applications, so only one install is required.

## 1.1 Starting the Interface

To run the CMM arm interface, as with any instrument interface in SpatialAnalyzer, the first step is to run SA and add the instrument model.  This is more than just a model; it supplies the connection for the instrument, and serves as the real-time representation of the instrument in SA.  Once the appropriate arm model has been added, run the arm interface.  We’ll use the Faro G08-05 as an example, so we’ll run Faro.exe.  The Spatial Analyzer CMM Log On Dialog appears with available SA connection shown:

In the case of the Faro interface, you need to select the serial port at which the arm is connected and the Baud rate for the arm connection.  As with all interfaces, you’ll need to type in your user name and select the desired SA connection from the SA Instrument Connection list (only one connection shown in this example).  As with most instrument interfaces, we need to tell the Log On if we are connecting to an arm.  An arm was not available at the time of this writing, so we’ll leave “Connect to Arm” unchecked, and run in simulation mode.  Now press  to connect to SpatialAnalyzer and the arm (simulation in this case).  The connections will be made, and the main SpatialAnalyzer CMM Interface dialog will appear:

Note the title bar includes “Sim.”, for simulation, in this instance.  This is a reminder that indeed we are not connected to an arm, and are running in simulation mode.  The title bar also contains the arm designation and the interface version number (release date).  Next to the SA logo, the interface displays the SpatialAnalyzer Instrument index.  We are connected to Instrument 0 in this example.  Note also the bottom row of buttons,

These buttons simulate the arm buttons when in simulation mode.  They are delineated by button number and duration of simulated button hit.  The label meanings are: 1S = button one, short duration; 1L = button one, long duration; 2S = button two, short duration; 2L = button two, long duration.  By determining the duration, or length of time the user holds down a button, the interface effectively adds an extra function to each button.  The SpatialAnalyzer CMM Interface allows you to program the action of each arm button for both short and long duration button hits.  IMPORTANT:  Since the interface uses each button to perform two functions by determining the duration which the button is held down, all actions are taken when an arm button is released, not when pressed.  So, you must remember when acquiring data that the data point is taken when you let go of the button, not when you press it.  See the next section for more information on programming the button modes.

## 1.2 Customizing the Interface

You can control many aspects of the behavior of the SpatialAnalyzer CMM Interface.  To change arm settings, press the  Arm Settings Button in the main dialog.  The CMM Interface Options Dialog will appear:

In the Button Functions section of the Options Dialog, you will be able to program the action of the arm buttons.  In the case of the Romer arm, there will appear a program interface for button three.  The default behavior for button one is “Record / Accept”, meaning that a short duration button hit (just press and let go) will result in a Record action, and a long duration (press and hold down for around 1.5 seconds) will result in an Accept action.  For most measurement modes, Record will acquire a single point, and Accept will accept the entire measurement.  The Accept function is the same as an OK button for an entire process.  The default behavior for button two is “Delete Last / Cancel”.  Again, the meaning is “Short duration action / Long duration action”.  For most measurement modes, Delete Last will remove the last point acquired, and Cancel will cancel an entire measurement.  To change the programming for  button two, press the  in the corresponding button function drop list.  You’ll see the available functions for button two in the drop list:

Just select a new program for the button by single left clicking the desired program in the list.  In this example, we’ll just keep the default behavior.  At the time of this writing, Delete Last / Mouse Toggle is supported only in the Romer interface.  When the Mouse mode is toggled on, the arm can be used in place of the mouse to move the Windows cursor around the screen, and the arm buttons can be used in place of the mouse buttons to perform general Windows actions.

allows you to set the sampling rate of the arm while it is acquiring position data.  In this case, the interface is set to query the arm 4 times per second during measurement operations.

allows you to tell the interface whether to send all measured points at the end of a geometry fit type of measurement.  For example, if checked, then at the end of a sphere fit measurement, the interface will not only send the fitted sphere and the sphere center point, it will also send all points which were actually measured with the arm and used for the fit.  If not checked, the interface will send only the sphere geometry and the sphere center point.

allows you to update the arm model in SpatialAnalyzer even when the arm is not measuring.  If left unchecked, the interface will update the model in SA only when actually acquiring data.  If you are working with an SA job which contains a large CAD file, you may want to leave this option unchecked in favor of better graphics performance.

is the setting for the frequency of SA Graphics updates when not measuring.  If  is not checked, then  does not apply.

The Probe section of the Options Dialog allows you to select the diameter of the probe currently in use.  In the case of the Romer interface, the probe diameter can be auto-detected through the interface.  To enter a new probe diameter, just type in the new diameter and press .  Whenever the probe is changed, a probe calibration should be performed.  More information on probe calibration is available in the section titled “Arm Calibration”.

The Material Thickness section of the Options Dialog allows you to set a material thickness for the Patch Measure Mode.  More information on material thickness will be presented in the section titled “Measure Modes”.

The Units section allows you to set the linear units displayed throughout the interface.  Just press the down arrow in the Units drop list to view/change the displayed units.

To view or change sound options, press  in the Options Dialog.  The Sound Properties Dialog will appear:

To enable Speech Prompting, you must be running the SASpeechServer.  Please see the SpatialAnalyzer User’s Manual for more information on the speech server.  If you are running the speech server, and would like to receive voice prompts, just check Speech Enabled, and press .  If your settings are correct, the SpatialAnalyzer voice will inform you that all is well.

The Arm Sounds section in the Sound Properties dialog allows you to customize the sounds the arm makes during certain operating procedures.  Max Error Proximity allows you to set the proximity at which an arm sound will be made.  This sound is meant to let you know the arm probe tip is headed away from the desired position.  “Good” Proximity is a range within which the arm will give another sound to let you know the arm probe tip is very near a desired position.  Use for Guided Measurements and Use for Plane Sections are the operating procedures to which the arm sounds are applied if checked.  More information on guided measurements and plane sections are presented in the section titled “Measure Modes”.  Press  to apply changes and close the Sound Properties dialog, or  to ignore changes and close the dialog.

Back in the Options Dialog, press  to view or change fit options.  The Fit Options Dialog will appear:

If, for example, Lock Sphere Radii is checked, then all fitted spheres will subsequently be locked to the radius you enter. This feature is commonly used when measuring tooling balls of known radii.  The fit routine is locked to the known radius, so that the only variables to solve for are the location of the center of the sphere.  Press  to accept and apply any changes you’ve made, or  to ignore changes.  The Fit Options dialog will close either way.

Once more, back in the Options Dialog, you’ll see the  button.  This button only appears when running in simulation mode as in this example.  Press  to get a look at the Joint Simulation Dialog:

This dialog allows you to update the arm model in SpatialAnalyzer when you are not connected to a real arm.  Since we are running a 7 degree of freedom arm in this example, there are 7 joints which can be manipulated.  Just click on the up or down arrow in one of the spin buttons  to move the corresponding joint in the SA model.  To change the angular Increment per spin button click, enter a new Increment, and press .  Press  or  to close the Joint Simulation Dialog.

Finally, back in the Options Dialog one last time, press  to apply changes and close, or  to ignore changes and close.

The next section will address the issue of arm calibration.  Following that, we’ll finally get to some measurements with the arm.

## 1.3 Arm Calibration

The SpatialAnalyzer CMM Interface allows for arm calibrations by calling the manufacture’s calibration routines.  Press  in the main Interface Dialog for a look at the arm calibrations.  In this case, we’ll get a connection failure notice since we are running in simulation mode, and the Faro Probe Calibrations dialog will appear:

The Arm Info section is not populated, again because we are in simulation mode, and we see some brief instructions for running the Faro Probe Calibration routines.  Press  for a look at the Faro Probe Cal interface:

Just enter the appropriate probe diameter (in mm as instructed in the previous dialog), and select a calibration from the Calibrate section.  The Faro interface will take you through a scripted process.

Once the arm is calibrated, you are ready to measure.  The next section will address that subject.

## 1.4  Measure Modes

Now that you’ve seen how to customize the arm interface and calibrate the arm, you’re ready to take some measurement data.  The Features section of the main SA CMM Interface dialog contains all the available measure modes:

You can hover over a measure mode button to get a mouse prompt text description of the mode.  For example, if you place the mouse over , a small text box next to the mouse cursor will say “Measure Single Points”.  To begin any measure mode, simply left click on the measure button.  Since we are running in simulation mode, all measure modes which provide a simulation will have a row of simulated arm buttons near the bottom of the dialog:

As described in the section titled “Starting the Interface”, These buttons simulate the arm buttons when in simulation mode.  They are delineated by button number and duration of simulated button hit.  The label meanings are: B1 Short = button one, short duration; B1 Long = button one, long duration; B2 Short = button two, short duration; B2 Long = button two, long duration.  And as described in the section titled “Customizing the Interface”, the buttons will perform the actions you set in the Button Functions section of the Options Dialog.  When running an arm, remember this IMPORTANT NOTE:  Since the interface uses each button to perform two functions by determining the duration which the button is held down, all actions are taken when an arm button is released, not when pressed.  So, you must remember when acquiring data that the data point is taken when you let go of the Record / Accept button, not when you press it.  This note does not apply to simulation mode, the button action is taken when the corresponding simulation button is pressed.  We’ll cover all the measure modes, corresponding arm button actions, and common usages next.  For all the mode descriptions, we’ll assume the default programming for the simulated arm buttons:

Single Points  Press .  The Points Measurement Dialog will appear:

Note the current Tip Diameter, set in the Options Dialog, is displayed.  Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm.  A data point will be acquired at the current position of the probe tip, and that point will be sent to SpatialAnalyzer.  The Points Measurement Dialog will now look similar to:

Press  again to acquire another point.  The dialog will now look like:

So the Points Measurement dialog always displays the Cartesian position of the two most recent points acquired.  Note that with this measure mode, points are sent to SA as soon as they are acquired, so the Delete button function takes no action.  To close the Points Measurement Dialog, press , a long duration of the Record / Accept button.  This is the equivalent an OK button.  In this case, the Cancel button function,  will perform the same function.  Since there is no process to cancel, it will simply close the dialog also.  Finally, the  button, which is available when connected to an arm, is a third equivalent way to close the Points Measurement Dialog.  This button is for users who may be at the computer screen and wish to close the dialog without having to press an arm button.  Most measure mode dialogs will have buttons similar to .

Button Functions:

- Record – acquires a single point.  The Point is immediately sent to SA.

- Accept – closes the dialog and ends the measurement.

- Delete Last – takes no action in this case.  All data is sent to SA as it is acquired.

- Cancel – closes the dialog and ends the measurement.

Common Uses:

Often measurement fixtures contain points which have relatively known positions.  These points are often used to locate a metrology instrument in the frame of the fixture.  The single Points Measurement mode is used to acquire individual points in discrete positions.

Stream Points:  Now press  in the main interface dialog.  The Stream Points Measurement Dialog will appear:

Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to begin streaming points to SpatialAnalyzer.

The interface will continue streaming points until you tell it to stop.  You can send separate streams by pausing and restarting the stream.  As instructed in the Stream Points Measurement Dialog, you pause with a short duration press of the Delete Last button.  You can subsequently restart a new stream by once again pressing the Record button.

Button Functions:

- Record – begins streaming points to SpatialAnalyzer.  The points are sent to SA as they are acquired.

- Accept – closes the dialog and ends the measurement.

- Delete Last – pauses streaming.

- Cancel – closes the dialog and ends the measurement.

Common Uses:

Streamed points are most often used to fit a geometry or a surface in SA.  The ability to pause and restart this mode is provided specifically for surface fits.  Any time data is to be fit with a method available in SpatialAnalyzer, Streamed Points is likely the best measure mode to use.

Patch Measurement:  Press  to measure Patch Surface Points.  The Surf Point Measurement Dialog will appear:

As usual, the Tip Diameter which we set earlier in the Options Dialog is shown.  Note also that “No Patch Shift” indicates the current “Material Thickness” setting from the Options Dialog.  As a reminder, press  to close the Surf Point Measurement Dialog.  Back in the main interface dialog, press  to go back to the Options Dialog and revisit the arm settings:

In the Material Thickness section, type in 0.05” (or an appropriate value), and check  to apply the Material Thickness.  Press  to accept the changes and close the Options Dialog.  Now press  in the main dialog again to go back into the Surf Point Measurement Dialog:

Now the dialog shows the 0.05” patch shift we just entered in the Options dialog.  Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the first plane (patch) point.  As with the single point mode, continue to press the Record button to acquire more plane points.  After at least three plane points have been acquired, the interface will fit a plane to all the plane points with each newly acquired point.  The results, ABCD of the fitted plane along with Maximum and RMS errors, will be displayed.  When you are done acquiring plane points, press  , or press and release a long duration (hold until you hear a beep, ~1.5 to 2 sec.) of the Record / Accept button if you’re running an arm, to accept the plane portion of the measurement.  You can also press the Accept Plane button, which becomes enabled as soon as at least 3 points are acquired.  Now press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the Projected Point.  The Surf Point Measurement Dialog will display the Cartesian location of the measured point.  If you are not satisfied with the measurement, just re-acquire.  If you are satisfied with the measurement, press  , or press and release a long duration (hold until you hear a beep, ~1.5 to 2 sec.) of the Record / Accept button if you’re running an arm, to accept the entire measurement.  You’ll also notice that the Accept Plane button has become the Accept Point button.  It has the same function – to accept the projected point, and indeed the entire measurement.

Button Functions:

- Record – acquires a single point.  If measuring the plane, the Record button also causes a plane fit to be performed so long as at least 3 plane points have been acquired.  If the plane has been accepted, acquires the projected point.

- Accept – If measuring the plane, accepts the fitted plane and makes the dialog ready for the Projected Point.  If the Projected Point has been measured, accepts the entire measurement, and closes the dialog.  The data is then sent to SA.  Note:  The long duration press of the Record / Accept button is functionally identical to the Accept button in the dialog.  The difference is that one button is located on the arm, the other is in the dialog.

- Delete Last – If measuring the plane, deletes the last measured plane point and eliminates it from subsequent plane fits.  If measuring the Projected Point, removes the last logged measurement.  You can also just Record over the last logged Projected Point.

- Cancel – closes the dialog and ends the measurement.  No data is sent to SA.  Note: The long duration press of the Delete Last / Cancel button is functionally identical to the Cancel button in the dialog.  The difference is that one button is located on the arm, and the other is in the dialog.

Common Uses:

The Patch Measurement, or Surface Point Measurement, is often used to measure a point on the “hidden” side of a surface.  For example, you may need to know the position of a point on the inside of the sheet metal skin of a machine part.  So long as the material thickness of the skin is known, the Patch Measurement mode can be used to measure the point.

Pin Measurement:  Press  to measure a pin.  The Pin Measurement Dialog will appear:

Note the Patch Shift is shown in this dialog.  A material thickness can be used in the same way as with the Patch Measurement.  Please see the Patch Measurement description in this section for more information on setting and using the Patch Shift.  This is another measure mode which contains more that one sub-procedure – a plane and a pin (circle) in this case.  Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the first plane (patch) point.  As with the single point mode, continue to press the Record button to acquire more plane points.  After at least three plane points have been acquired, the interface will fit a plane to all the plane points with each newly acquired point.  The results, ABCD of the fitted plane along with Maximum and RMS errors, will be displayed.  When you are done acquiring plane points, press  , or press and release a long duration (hold until you hear a beep, ~1.5 to 2 sec.) of the Record / Accept button if you’re running an arm, to accept the plane portion of the measurement.  You can also press the Accept Plane button, which becomes enabled as soon as at least 3 points are acquired.  Now press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the first Pin point.  After at least 3 Pin points have been acquired, the interface will fit a circle to all the pin points with each newly acquired point.  The Radius, Maximum error, and RMS error of the Pin (circle) will be displayed.  If you are satisfied with the measurement, press  , or press and release a long duration (hold until you hear a beep, ~1.5 to 2 sec.) of the Record / Accept button if you’re running an arm, to accept the entire measurement.  You’ll also notice that the Accept Plane button has become the Accept Pin button.  It has the same function – to accept the pin measurement, and close the dialog.

Button Functions:

- Record – acquires a single point.  If measuring the plane, the Record button also causes a plane fit to be performed so long as at least 3 plane points have been acquired.  If measuring the pin, the Record button also causes a circle fit to be performed so long as at least 3 circle points have been acquired.

- Accept – If measuring the plane, accepts the fitted plane and makes the dialog ready for the Pin (circle) points.  If the Pin Points have been measured, projects the pin circle points to the plane, accepts the entire measurement, and closes the dialog.  The data is then sent to SA.  Note:  The long duration press of the Record / Accept button is functionally identical to the Accept button in the dialog.  The difference is that one button is located on the arm, the other is in the dialog.

- Delete Last – If measuring the plane, deletes the last measured plane point and eliminates it from subsequent plane fits.  If measuring the pin, deletes the last measured pin point and eliminates it from subsequent circle fits.

- Cancel – closes the dialog and ends the measurement.  No data is sent to SA.  Note: The long duration press of the Delete Last / Cancel button is functionally identical to the Cancel button in the dialog.  The difference is that one button is located on the arm, the other is in the dialog.

Common Uses:

Pins, like tooling balls, are often used to locate instruments with respect to CAD models.  The Pin mode also serves as a cylinder measurement.  Given that the resulting circle is projected to the plane, and can be shifted by the arm tool tip diameter plus a patch offset, the measurement result can be located at the true base of the cylinder, or even inside the planar surface.

Hole Measurement:  Press  to measure a hole.  The Hole Measurement Dialog will appear:

This measurement mode is functionally identical to the Pin Measurement.  The only difference is that the pin measurement involves fitting a circle to the base of a cylindrical object and projecting the fitted circle to a previously measured plane, whereas the Hole Measurement involves measuring a circular hole and projecting that to a previously measured plane.  Please see the description of the Pin Measurement in this section for details.

Slot Measurement:  Press  to measure a slot.  The Slot Measurement Dialog will appear:

The Slot Measurement is procedurally quite similar to the Pin and Hole Measurements.  The difference is the fact that the Slot Measurement involves two holes (circles) instead of one.  Therefore, there is simply one extra step in the Slot Measurement.  Please see the description of the Pin Measurement in this section for details.  When the Slot Measurement is done and ready for acceptance, the Slot Measurement will look similar to:

Note the extra results displayed near the bottom of the Slot Measurement Dialog.  These report the width, length, and center of the slot after both holes (Circle A and Circle B) have been fitted and projected to the Plane.  Recall that in all cases, the B1 Short, etc. buttons are only shown when running in simulation mode, without an arm connected.

Line Measurement:  Press  to measure a line.  The Line Measurement Dialog will appear:

Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the first line point.  The dialog will then look something like:

The Cartesian position of the first point, P1, is shown.  Now press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, once more to acquire the second line point.  The dialog will then look something like:

Now both line points, P1 and P2, are displayed.  Delta is the Cartesian difference in the point positions, and Length is the resulting length of the line.

Button Functions:

- Record – acquires a line point.

- Accept – accepts the line, closes the dialog, and sends the data to SA.  Does nothing if both line points have not been measured.

- Delete Last – removes the last measured line point.

- Cancel – closes the dialog and ends the measurement.  No data is sent to SA.

Circle Measurement:  Press  to measure a circle.  The Circle Measurement Dialog will appear:

The Circle Measurement procedure is nearly identical to the circle measurement steps in the Pin, Hole, and Slot Measurements.  Note that you are required to set the radial measurement offset .  This is because the general circle fit can offset the fitted circle with respect to the measured data, but must be told which direction to offset.  In the Pin Measurement, the circle is always measured Outside the Radius.  In the Hole and Slot Measurements, the circles are always measured Inside the Radius.  Please see the description of the Pin Measurement in this section for all other details regarding the Circle Measurement procedure.

Plane Measurement:  Press  to measure a plane.  The Plane Measurement Dialog will appear:

The Plane Measurement is functionally identical to the first step in the Pin, Hole, and Slot Measurements.  Please see the description of the Pin Measurement in this section for details.

Sphere Measurement:  Press  to measure a sphere.  The Sphere Measurement Dialog will appear:

Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to acquire the first sphere point.  As with the single point mode, continue to press the Record button to acquire more sphere points.  After at least four sphere points have been acquired, the interface will fit a sphere to all the sphere points with each newly acquired point.  The Sphere Measurement Dialog will then resemble:

The results displayed are the Cartesian location of the fitted Center of the sphere, the resulting Radius of the sphere, the Percent Coverage of the sphere given the fit and the locations of the measured points, and the RMS Error of the fit.

Button Functions:

- Record – acquires a single point.  If at least 4 points have already been acquired, the Record button also causes a sphere fit to be performed.

- Accept – Accepts the entire measurement, and closes the dialog.  The data is then sent to SA.  Note:  The long duration press of the Record / Accept button is functionally identical to the Accept Sphere button in the dialog.  The difference is that one button is located on the arm, the other is in the dialog.

- Delete Last – Deletes the last measured sphere point and eliminates it from subsequent sphere fits.

- Cancel – closes the dialog and ends the measurement.  No data is sent to SA.  Note: The long duration press of the Delete Last / Cancel button is functionally identical to the Cancel button in the dialog.  The difference is that one button is located on the arm, and the other is in the dialog.

Cross Section Measurement:  Press  to measure a Cross Section.  The Section Specification Dialog will appear:

The Section Location section tells the Cross Section Measurement where to collect data.  To use this measure mode, you must be familiar with creating frames and setting the Working Frame in SpatialAnalyzer.  Please see the SpatialAnalyzer User’s Manual for more information.  Be sure the appropriate frame in SA is the current Working Frame.  Select the appropriate plane in the Section Location area.  For example, if Z is selected, then a plane parallel to the XY plane will be defined.  Type in the value for the plane.  For example, -18.28 in. defines a plane parallel to the XY plane, and crossing the Z-axis at Z = -18.28 inches.  Press OK to accept the Section Specification.  The Section Measurement Dialog will appear:

Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to begin sampling data.  A new point will be acquired and sent to SpatialAnalyzer whenever the center of the arm probe tip crosses the plane you defined in the Section Specification Dialog.  The interface will continue sampling data until you tell it to stop.  As instructed in the Section Measurement Dialog, you can pause sampling with a short duration press of the Delete Last button.  You can subsequently restart a new sample stream by once again pressing the Record button.

Button Functions:

- Record – begins sampling data.  Points are acquired and sent to SA when the center of the arm probe tip crosses the plane defined in the Section Specification Dialog.

- Accept – closes the dialog and ends the measurement.

- Delete Last – takes no action in this case.  All data is sent to SA as it is acquired.

- Cancel – closes the dialog and ends the measurement.

Common Uses:

Section Measurement could be considered the inverse of a plane fit.  Section Measurement allows you to define a plane in space, and then acquire data which lies on that plane.

Frame Measurement:  Press  to measure frames.  The Frame Measurement Dialog will appear:

Type in a name for the first frame, and press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm.  A frame will be acquired with the origin at the current position of the probe tip, and the orientation such that the Z-axis of the frame is parallel with the axis of the probe tip.  The frame will be sent to SpatialAnalyzer, with the Name you have given it.  The name in the dialog will be incremented, and the Frame Measurement Dialog will now look similar to:

So far only one frame has been recorded, so Current and Last show the same frame.  The reported results are the translation (xyz in current display units), and rotation (degrees) of the measured frame with respect to the current Working Frame in SA.

Button Functions:

- Record – acquires a Frame.  The Frame is immediately sent to SA.

- Accept – closes the dialog and ends the measurement.

- Delete Last – takes no action in this case.  All data is sent to SA as it is acquired.

- Cancel – closes the dialog and ends the measurement.

Common Uses:

Frame Measurement is the same as point measurement with the addition of orientation.  Point Measurements result in three dimensional locations, whereas Frame Measurements result in six dimensional positions and orientations.

Guided Measurement:  To use this mode, you must be familiar with Batch Point Measurement in SpatialAnalyzer.  In SA, go to the Instrument menu, and expand Automatic Measurement.  Select Measure Batch of Points.  You’ll be prompted to select the points you’d like to measure, and for a group and target suffix for the batch.  You will then be prompted to select the instrument for transmitting the batch.  When the batch instructions are transmitted, the Guided Measurement button will become enabled. Press  to begin the Guided Measurement.  The Guided Measurement Dialog will appear:

Note the absence of the B1 Short, etc. buttons.  The Guided Measurement does not have a simulation mode.  The dialog shows the identifying Group and Target names of the current Nominal point in the batch.  The Cartesian position of the current Nominal point is shown.  The real-time Actual Cartesian position of the arm probe tip is shown.  The Cartesian Delta between the batch point and the probe tip is shown, and the magnitude Deviation between the batch point and the probe tip is shown.  Press , or press and release a short duration (<=1 sec.) of the Record / Accept button if you’re running an arm, to accept the batch point and move on to the next one.

Button Functions:

- Record – accepts a single point.  Same as pressing the Accept Point button.

- Accept – closes the dialog and ends the measurement.  Same as pressing the Close button.

- Delete Last – takes no action in this case.  All data is sent to SA as it is acquired.

- Cancel – closes the dialog and ends the measurement.  .  Same as pressing the Close button.

Common Usage:

The Batch Measurement can be used as a diagnostic to check the drift of a fixture with respect to the instrument.

Inspections:  Notice  in the main dialog.  Any measure button which includes the  in its icon can be used to inspect a previously measured entity of the same type.  For example,  right press  to inspect a previously measured pin.  The Pin Measurement Dialog will not appear immediately.  First, drag the mouse over SpatialAnalyzer.  You’ll see a mouse prompt instructing you to “Select Circle to Inspect”.  Select the Pin in question by left double clicking directly on the representative circle in SA.  Now the Pin Inspection Dialog will appear:

Note the dialog title bar indicates the name of the Pin Center in SA.  Measure the pin exactly as described in the Pin Measurement description earlier in this section.  When the pin is accepted, the Comparison Results Dialog will appear:

Given the nature of the simulation mode, our results are absolutely horrible.  Nonetheless, the Comparison Results Dialog presents the Pin Inspection results.  Press  to close the Comparison Results and Pin Inspection Dialogs.  The Hole and Plane inspections are performed in the same way as the Pin Inspection.